[image: image1.png]



DAILY VACUUM PUMP INSTRUCTIONS

Turn on and warm up the pump and Peltronic condenser for 15 minutes before use.  

Check light on Peltronic condenser has turned green before use.

Run air through the outlet for 30 seconds after each run.

Every evening, turn off the pump and EMPTY the catchpots.

NEVER leave solvent in the pots when the pump is not running!

WEEKLY PUMP INSTRUCTIONS

As above, plus:

Open the vac line and purge with air through the pump for 10 minutes.

Follow the above steps before a prolonged shutdown, e.g. Xmas holiday.

MONTHLY PUMP INSTRUCTIONS

As both steps plus:

Shut gas ballast cap.  Check ultimate vacuum of pump with accurate vacuum gauge (1mbar resolution or better).  Make a note of this reading for regular comparison against the ultimate vacuum of your model, (shown on the serial number plate).

If the pump is more than 15 mbar above the ultimate vacuum inform a Technician.  Do not use the pump until remedial action has been taken.  Inform a technician if the pump becomes noisy or you have any other concerns about its operation.
VACUUBRAND CONTACT – GEORGINA CAPE 07974 028853

