General Risk Assessment Form

| Date: (1) 30/08/2017 | Assessed by: (2) | Checked / Validated* by: (3) | Location: (4) | Assessment ref no: (5) Chromatography Equipment | Review date: (6) |
| --- | --- | --- | --- | --- | --- |
| Task/premises: (7) - The use of Gas Chromatography and High Performance Liquid Chromatography with UV –VIS, Flame ionization detection and Mass Spectrometry. |  |  |  |  |  |

| Activity (8) | Hazard (9) | Who might be harmed and how (10) | Existing measures to control risk (11) | Risk rating (12) | The result (13) |
| --- | --- | --- | --- | --- | --- |
| Routine Use of GC-MS | electrical shock | User / burns and/ or death | Equipment is built to EU standards for electrical safety and EU mark is displayed. Equipment is maintained correctly as per user manual and manufactures recommendations [Lab Group to insert planned preventative maintenance schedule as an annex to this risk assessment.] Equipment is sourced from reputable supplier and / or manufacturer Equipment is PAT tested annually. The electrical wiring is kept away from water and sources of heat which could degrade wirings performance. Instrument is not used when panels or outer casing is removed. Correct training is given for the use of the equipment as per user manual. | 1*3 =3 | A |
| Routine Use of GC-MS | Asphyxiation | User and others in laboratory / death or brain damage | No more than the minimum necessary bottles of carrier gas is stored in the room. All regulators must be changed as per the manufactures guidance for the gas type. Use the calculation below to calculate if it is safe to store the Gas cylinder in the room. If the calculation = < 19% contact the Health and Safety office for advice. Volume of oxygen= 0.205 X the volume of the room- volume of inert gas liberated [If > 19% then the cylinder is safe to use in this room without monitoring.] [If < 19% then the room will need oxygen monitoring.] | 1*3 = 3 | A |
| Routine Use of GC-MS | oil fumes | User / Chemical exposure and Smoke inhalation | Correct maintenance carried out by competent person. PPM followed as per the manufacturers recommendations. Pumps must be serviced correctly by a trained individual. | 1*2 = 3 | A |
| Routine Use of GC | fire | User and others in building / Smoke inhalation, burns, or death | GC, MS and other hot parts must cited well away from solvents and other flammable material. H2 Cylinders are not to be used. Generators are to be used instead. [If H2 cylinders have to be used group is to insert reasons here] Flash back arrestor is to be used on any H2 line. Fire detection system in room to warn of fire. Appropriate firefighting medium in room for type of fire. [Group to decide type of extinguishing medium] | 1*3 =3 | A |
| Routine Use of GC | damage to hearing | User and others in laboratory / WRHL | If noise levels become too high contact the Safety Office for an official test of noise levels and a remedial plan. Site vacuum pumps and compressors in vented cupboards. | 1*1 = 1 | T |
| Routine Use of GC | entrapment of digits | User / Minor injury to finger | Do not interfere with the auto sampler when it is in motion. Keep clear. Sign is to be put in place to warn of danger if instrument is sited in a multiple supervised area. | 1*1 =1 | T |
| Routine Use of GC | Thermal burns | User / minor to medium burns to digits | Parts of the GC and MS are HOT allow sufficient time for them to cool before being used or touched. Sign is to be put in place to warn of danger if instrument is sited in a multiple supervised area. | 1*1 = 1 | T |
| Changing columns | Sharps | User / Minor injury to finger | When cutting columns only used approved cutting tools and methods. The following items of PPE must be worn: Flame retardant laboratory coat BS EN374 compliant gloves (nitrile) BS EN166 compliant eye protection (chemical splash proof safety glasses). Users must be trained before carrying out this task. | 1*1 = 1 | T |
| Changing GC columns | Thermal burns | User / minor to medium burns to digits | Parts of the GC and MS are HOT allow sufficient time for them to cool before being used or touched. Sign is to be put in place to warn of danger if instrument is sited in a multiple supervised area. | 1*1 = 1 | T |
| Routine use of HPLC | Electrical Shock | User / burns and/ or death | Equipment is built to EU standards for electrical safety and EU mark is displayed. Equipment is maintained correctly as per user manual and manufactures recommendations [Lab Group to insert planned preventative maintenance schooled as annex to this risk assessment.] Equipment is sourced from reputable supplier and / or manufacturer Equipment is PAT tested annually. Equipment is plugged into a RCD controlled socket. The electrical wiring is kept away from water and sources of heat which could degrade wirings performance. Instrument is not used when panels or outer casing is removed. Correct training is given for the use of the equipment as per user manual. | 1*3 = 3 | A |
| Routine use of HPLC | Fire | User and others in building / Smoke inhalation, burns, or death | Use of flammable solvents. Only minimum quanties of these substances are to be used. Solvent reservoirs are kept away from source of ignition i.e. naked flames Equipment is correctly maintained as per user manual. Fire detection system in room to warn of fire. Appropriate firefighting medium in room for type of fire. [Group to insert type of firefighting medium here] | 1*3 = 3 | A |
| Routine use of HPLC | Chemical Contamination | User / See CRA for details | The following items of PPE must be worn: Flame retardant laboratory coat BS EN374 compliant gloves (nitrile) BS EN166 compliant eye protection (chemical splash proof safety glasses). A Chemical Risk Assessment must be completed for the solvents and salts used in the mobile phase. User must include this risk assessment with their CRA to cover the use of the analytes. Solvents must not be poured at the top at the top of the HPLC tack. Reservoirs must be brought to bench level. | See CRA | A |
| Routine use of HPLC – UV-VIS | Non ionizing radiation (200nm – 700nm) | User / Burns and eye damage | Source is sealed within equipment no exposure when in use. Interlock in place to stop experiment if source is exposed. | 1*2=2 | A |

| Action plan (14) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Ref No | Further action required | Action by whom | Action by when | Done |
|  | This form will be altered to suit each laboratory groups specific operating procedures. | Lab group |  |  |
|  | Lab group is to implement and maintain LabCup asset register. Reviewed annually at inspection | Lab group / H+S Office | On-going |  |
|  | Lab group to create a PPM for the instruments they have and attach plan to this Risk Assessment, | Lab group |  |  |
|  | Groups are to create risk assessments for the maintenance of the instruments | Lab Groups |  |  |

| Signed and understood. By signing below you are acknowledging that you have read the risk assessment and understand the risks involved and the control measures in place to protect you and other users of the laboratory. |  |  |
| --- | --- | --- |
| Name | Signature | Date |